## information
article title | Standardization of radiation therapy quality control system through mutual quality control based on failure mode and effects analysis
journal name | Radiological Physics and Technology
author names | Yuki Tanimoto | Masataka Oita | Kazunobu Koshi | Kiyoshi Ishiwaki | Futoshi Hiramatsu | Toshihisa Sasaki | Hiroki Ise | Takashi Miyagawa | Takeshi Maeda | Shinsuke Okahira | Takashi Hamaguchi | Tatsuya Kawaguchi | Norihiro Funada | Shuhei Yamamoto | Akira Hiroshige | Yuki Mukai | Shohei Yoshida | Yoshiki Fujita | Atsuki Nakahira | Hirofumi Honda
corresponding author | Masataka Oita, Ph.D.
affiliation | Graduate School of Interdisciplinary Science and Engineering in Health Systems, Okayama University
e-mail address | oita-m@cc.okayama-u.ac.jp

## Online Resource 1
Online Resource 1 Identification of high-risk factors by FMEA
Process | Failure Mode | 1 or fewer qualified personnel per linear accelerator | [7 empty] | 2 or more qualified personnel per linear accelerator
|  | Occurrence (O) |  | Severity (S) |  | Detectability (D) |  | Risk Priority Number (RPN) |  | Occurrence (O) |  | Severity (S) |  | Detectability (D) |  | Risk Priority Number (RPN)
|  | Mean | SD | Mean | SD | Mean | SD | Mean | SD | Mean | SD | Mean | SD | Mean | SD | Mean | SD
Acceptance test | Check items were overlooked and not implemented. | 2.33 | 0.82 | 3.5 | 1.05 | 4.65 | 0.16432 | 38.9 | 15.81 | 2.43 | 0.79 | 4 | 1.83 | 4.67 | 0.16036 | 44.23 | 19.69
Acceptance test | Equipment used for verification was malfunctioning, and accurate values could not be obtained. | 2.83 | 1.33 | 3.5 | 0.84 | 6.6 | 0.10954 | 69.48 | 39.58 | 2.43 | 0.53 | 3.86 | 1.07 | 6.61 | 0.1069 | 61.19 | 17.26
Commissioning of radiation detection equipment | The detector was calibrated at a different voltage than that used in their institution. | 2.33 | 1.03 | 4.33 | 1.86 | 4.75 | 0.054772 | 51.7 | 36.14 | 2 | 1 | 3.71 | 0.49 | 4.76 | 0.053452 | 36.06 | 20.37
Commissioning of radiation detection equipment | Uncertain measurements were obtained due to a lack of understanding of the instrumental characteristics required to obtain stable values. | 4.17 | 1.47 | 3.67 | 1.21 | 5.6 | 0.10954 | 87.25 | 46.95 | 3.14 | 0.69 | 4.43 | 1.13 | 5.61 | 0.1069 | 77.96 | 28.49
Beam modeling | The direction of operation of the detector used to measure beam data was moved in the direction of greater uncertainty. | 2.67 | 1.37 | 5 | 1.26 | 4.55 | 0.27386 | 62.85 | 39.89 | 1.86 | 0.9 | 4.29 | 1.38 | 4.59 | 0.26726 | 38.74 | 25
Beam modeling | Inappropriate detector was selected for small field measurement. | 2.67 | 1.75 | 5.17 | 1.72 | 5.2 | 1.6432 | 66.63 | 39.15 | 2 | 0.82 | 3.71 | 0.76 | 5.41 | 1.6036 | 41.4 | 28.3
Beam modeling | Water in the water phantom evaporated owing to the long measurement time, and the measurement depth was shifted. | 5.17 | 2.48 | 4.83 | 1.83 | 4.05 | 0.27386 | 115.27 | 95.92 | 4.29 | 1.25 | 4.43 | 1.72 | 4.09 | 0.26726 | 85.03 | 61.53
Beam modeling | The reference detector was not used, so the uncertainty of the field detector could not be corrected. | 3.5 | 1.87 | 4.5 | 0.55 | 3.85 | 0.16432 | 59.98 | 32.88 | 2.29 | 1.11 | 4 | 0.82 | 3.87 | 0.16036 | 37.43 | 24.51
Beam modeling | Beam data from another linac were registered during registration to the treatment planning system. | 1.83 | 0.98 | 3.17 | 1.47 | 8.6 | 0.10954 | 54.77 | 45.75 | 1.14 | 0.38 | 3.29 | 0.76 | 8.61 | 0.1069 | 32.03 | 10.94
Beam modeling | 3D water phantom was not properly set up and was measured in the wrong position | 4.5 | 1.87 | 4.5 | 1.64 | 6.95 | 0.27386 | 155.73 | 96.77 | 3.29 | 0.95 | 3.86 | 0.69 | 6.99 | 0.26726 | 91.77 | 43.58
Beam modeling | The results for the 3D water phantom were not acceptable; therefore, they were validated using a solid phantom. | 2.17 | 0.75 | 4.5 | 1.22 | 5.15 | 0.16432 | 50.03 | 19.64 | 1.57 | 0.79 | 4.29 | 0.95 | 5.17 | 0.16036 | 38.1 | 28.87
Determination of commissioning baseline | Uncertain values were established because of the short duration in which the baseline was obtained. | 4.33 | 1.21 | 4 | 1.1 | 5.05 | 0.27386 | 93.42 | 53.92 | 3.29 | 0.76 | 3.29 | 0.49 | 5.09 | 0.26726 | 54.21 | 12.83
Determination of commissioning baseline | Equipment was installed in the wrong position due to ambiguous confirmation of installation of the SSD/SAD to be measured. | 3.83 | 1.47 | 4 | 1.1 | 6.25 | 0.27386 | 96.17 | 44.36 | 2.57 | 0.98 | 4 | 0.82 | 6.29 | 0.26726 | 64.57 | 24.77
Commissioning of linac | Dosimetric leaf gap values were not optimal. | 3.67 | 1.21 | 3.5 | 1.05 | 6.2 | 1.0954 | 87.33 | 60.28 | 3.29 | 1.38 | 3.14 | 0.69 | 6.34 | 1.069 | 63.83 | 27.3
Commissioning of linac | Mistakes in entering wedge factor into the treatment planning system. | 2.5 | 0.55 | 3.83 | 1.17 | 6.4 | 2.3004 | 58.33 | 23.96 | 1.71 | 0.49 | 3.57 | 1.27 | 6.1 | 2.245 | 37.21 | 20.32
Commissioning of QA equipment | Sensitivity of the detector was reduced because the detector was not calibrated. | 4.33 | 1.86 | 4.33 | 0.52 | 5.55 | 0.27386 | 101.63 | 35.33 | 3.43 | 1.51 | 3.57 | 0.53 | 5.59 | 0.26726 | 67.83 | 30.88
Commissioning of QA equipment | Insufficient explanation of the measurement procedure, and continued use of incorrect measurement method. | 4.67 | 1.97 | 4.33 | 1.37 | 7.15 | 0.16432 | 148.27 | 80.01 | 3.71 | 1.38 | 3.86 | 1.07 | 7.17 | 0.16036 | 109.7 | 73.54
Commissioning of QA equipment | Angle meter was improperly adjusted and measured at the wrong angle. | 3 | 1.67 | 4.83 | 1.33 | 6.1 | 0.10954 | 95.1 | 83.42 | 2.57 | 0.53 | 3.86 | 1.07 | 6.11 | 0.1069 | 61.2 | 24.63
Commissioning of QA equipment | Embedded temperature/barometer malfunctioned and measured values incorrectly. | 3.17 | 0.98 | 4.83 | 1.33 | 5.75 | 0.27386 | 90.83 | 51.63 | 2.29 | 0.49 | 3.57 | 0.98 | 5.71 | 0.26726 | 46.21 | 13.87
Commissioning of QA equipment | Uncertain values were set owing to the short duration for obtaining the baseline. | 4.17 | 1.72 | 4.67 | 1.21 | 5.35 | 0.16432 | 104.85 | 47.39 | 3.14 | 0.9 | 3.71 | 0.76 | 5.37 | 0.16036 | 61.93 | 19.03
Determination of QA frequency and baseline | Required QC not itemized and not performed. | 4.67 | 1.86 | 3.83 | 0.75 | 5.75 | 0.27386 | 108.25 | 59.1 | 4 | 0.58 | 3.86 | 1.46 | 5.79 | 0.26726 | 92.86 | 53.03
Determination of QA frequency and baseline | Standard SSD/SAD settings were vague. | 4.17 | 1.47 | 4 | 1.55 | 5.5 | 0.21909 | 97.55 | 71.81 | 2.57 | 0.79 | 3.43 | 0.79 | 5.53 | 0.21381 | 48.43 | 19.17
Determination of QA frequency and baseline | Required QC was not included in the item. | 4.33 | 2.07 | 4.5 | 1.22 | 1.8 | 0 | 36.6 | 23.49 | 3.43 | 1.51 | 3.86 | 0.69 | 1.8 | 2.3984e-16 | 24.17 | 12.25
Preparation of measurement sheets | Incorrect spreadsheet citations prevented accurate calculations. | 5.67 | 1.86 | 5.17 | 1.72 | 7.9 | 0.10954 | 237.83 | 115.16 | 4.57 | 2.23 | 3.86 | 1.07 | 7.91 | 0.1069 | 147.14 | 110.98
Preparation of measurement sheets | A spreadsheet with incorrect polarity notation was used. | 4.83 | 1.94 | 5.33 | 1.63 | 4.45 | 0.27386 | 108.7 | 32.24 | 3.17 | 1.6 | 4.43 | 0.98 | 4.49 | 0.26726 | 53.49 | 41.43
Preparation of measurement sheets | Incorrect values were entered into spreadsheet formulas, resulting in erroneous calculations. | 5.33 | 1.21 | 5.5 | 1.76 | 7.47 | 0.40825 | 227.65 | 109.33 | 5.29 | 1.5 | 4.57 | 1.51 | 7.3 | 9.5934e-16 | 186.67 | 110.14
Preparation of measurement sheets | Incorrect results were obtained due to lack of uniformity in the usage of the unit of pressure. | 3.5 | 1.38 | 5 | 2 | 5.27 | 0.37238 | 86.93 | 30.49 | 3.14 | 1.07 | 4.43 | 1.27 | 5.11 | 0.1069 | 71.54 | 32.6
Preparation of measurement sheets | Different tests were performed because the names of the tests did not conform to the guidelines. | 3.83 | 1.6 | 4.5 | 1.22 | 4.8 | 0 | 80.8 | 31.32 | 2.86 | 0.69 | 3.71 | 0.49 | 4.8 | 0 | 52.11 | 17.4
Preparation of measurement sheets | Electrometer values were manually entered into the spreadsheet incorrectly. | 5.33 | 2.25 | 5.17 | 1.6 | 5.92 | 0.38687 | 155.42 | 53.53 | 4.14 | 1.86 | 4.29 | 1.25 | 5.76 | 0.053452 | 101.83 | 56.05
Safety management system QA | Collision interlock did not operate properly and struck the patient | 2.83 | 0.75 | 6.33 | 2.25 | 8.87 | 0.40825 | 158.27 | 66.79 | 2.14 | 0.38 | 4.71 | 1.38 | 8.7 | 1.9187e-15 | 88.24 | 30.32
Safety management system QA | Room had to be enterd during irradiation because the indicator light signaling that irradiation was in progress malfunctioned. | 1.5 | 0.55 | 5.83 | 2.32 | 4.7 | 0 | 37.6 | 8.92 | 1.29 | 0.49 | 4.43 | 1.72 | 4.7 | 0 | 28.2 | 17.79
Safety management system QA | Irradiation could not be performed because debris accumulating in the door groove prevented its closure. | 1.33 | 0.52 | 6.67 | 2.94 | 3.25 | 0.054772 | 27.17 | 12.2 | 1.43 | 0.79 | 4.14 | 1.95 | 3.24 | 0.053452 | 18.07 | 8.9
Safety management system QA | During irradiation, a person in the blind spot of the surveillance monitor was irradiated accidentally. | 2.5 | 1.22 | 6 | 2.76 | 8.3 | 1.9459e-15 | 121.73 | 84.26 | 2 | 1.41 | 4.14 | 1.07 | 8.3 | 1.9187e-15 | 67.59 | 50.36
Safety management system QA | Door interlock could not be unlocked; hence, irradiation could not be performed. | 1.83 | 0.75 | 7.33 | 2.5 | 3.3 | 4.8648e-16 | 47.3 | 31.07 | 2 | 0.82 | 4.86 | 2.12 | 3.3 | 4.7967e-16 | 31.11 | 17.03
Safety management system QA | Microphone had deteriorated, and the patient's voice and buzzer became inaudible. | 2.5 | 0.55 | 7.33 | 2.07 | 4.7 | 0 | 86.17 | 28.92 | 3 | 1.15 | 5.43 | 2.15 | 4.7 | 0 | 74.53 | 33.39
Mechanical QA | MLC stop position was within tolerance for static, but beyond tolerance for dynamic. | 3.67 | 1.63 | 5.17 | 1.47 | 4.9 | 2.0813 | 97.17 | 75.74 | 3 | 1 | 4 | 1 | 5.17 | 2.0312 | 65.2 | 45.27
Mechanical QA | Despite the accuracy of the laser having decreased after an earthquake, irradiation could be performed at the matched location. | 2 | 0.63 | 4.83 | 1.17 | 4.4 | 0.10954 | 44.82 | 26 | 1.86 | 0.69 | 3.67 | 0.52 | 4.41 | 0.1069 | 24.73 | 14.07
Mechanical QA | Laser position was acceptable for NonIMRT but not for IMRT. | 2.67 | 1.21 | 5 | 1.67 | 3.6 | 0.9859 | 57.3 | 50.76 | 2.43 | 0.53 | 3.57 | 0.53 | 3.73 | 0.96214 | 30.99 | 6.32
Mechanical QA | Mirror was misaligned and light irradiation fields were projected on to different positions. | 2.5 | 1.05 | 5.17 | 0.98 | 3.3 | 0.54772 | 43 | 20.36 | 2.29 | 0.95 | 4.14 | 1.21 | 3.23 | 0.53452 | 28.4 | 9.08
Dosimetry QA | Output calibration did not include required coefficients. | 3.17 | 1.47 | 5.17 | 1.47 | 7.8 | 0 | 132.6 | 82.55 | 2.57 | 1.4 | 4.14 | 1.07 | 7.8 | 9.5934e-16 | 92.49 | 74.87
Dosimetry QA | Correction factors had been fluctuating but had not been checked since purchase. Therefore, they were corrected with incorrect values. | 3 | 1.79 | 4.83 | 1.17 | 5.9 | 0.10954 | 93.3 | 82.14 | 2.29 | 0.76 | 3.86 | 0.69 | 5.91 | 0.1069 | 52.57 | 22.34
Dosimetry QA | Inaccurate measurements were obtained due to age-related deterioration of measuring equipment. | 3.67 | 0.82 | 5.33 | 1.51 | 5.65 | 0.16432 | 114.62 | 52.93 | 3.14 | 0.9 | 3.71 | 0.76 | 5.67 | 0.16036 | 67.54 | 25.86
Dosimetry QA | Output dose calibration did not include required factors. | 3.67 | 1.37 | 4.83 | 1.17 | 7.15 | 0.16432 | 123.42 | 41.53 | 3 | 1.15 | 4 | 1.41 | 7.17 | 0.16036 | 93.79 | 74.09
Dosimetry QA | Factors were uncertain because accurate temperature and pressure measurements could not be taken due to aging of the equipment. | 3.5 | 1.05 | 4.67 | 1.37 | 5.92 | 0.38687 | 100.53 | 55.97 | 3.14 | 0.9 | 3.71 | 0.49 | 5.76 | 0.053452 | 67.43 | 21.45
Dosimetry QA | The spreadsheet used for proofreading was from a different energy | 3.5 | 0.84 | 5.67 | 1.63 | 7.25 | 0.054772 | 141.6 | 43.39 | 3.71 | 1.5 | 4.57 | 1.13 | 7.26 | 0.053452 | 130.57 | 79.59
Dosimetry QA | Integral calibration was performed, but measurements were taken with different electrometers in different combinations. | 2.5 | 0.55 | 5.83 | 1.33 | 4.9 | 0.10954 | 74.9 | 33.16 | 1.43 | 0.79 | 4.29 | 1.6 | 4.91 | 0.1069 | 29.54 | 16.86
Dosimetry QA | Measurements were taken with the polarity of the potential applied to the electrometer reversed. | 2.33 | 0.52 | 5.33 | 1.86 | 4.25 | 0.054772 | 55.43 | 31.76 | 2.14 | 1.07 | 4.43 | 1.51 | 4.26 | 0.053452 | 43.24 | 29.85
Dosimetry QA | Measurements were conducted without setting a reference field. | 2.83 | 1.17 | 5.5 | 1.64 | 6.75 | 0.054772 | 111.75 | 74.04 | 3 | 1.29 | 4.57 | 1.13 | 6.76 | 0.053452 | 98.46 | 57.68
Dosimetry QA | Measurements were performed without setting a reference irradiation field. | 3.5 | 1.38 | 4.83 | 1.17 | 6.4 | 0.10954 | 111.3 | 65.36 | 3.14 | 1.07 | 4.43 | 1.51 | 6.41 | 0.1069 | 91.43 | 46.3
Dosimetry QA | Measurements conducted without changing the coefficients after calibration yielded incorrect results. | 3 | 1.41 | 4.5 | 1.22 | 4.9 | 0.10954 | 66.7 | 34.33 | 2.57 | 1.27 | 3.29 | 0.76 | 4.91 | 0.1069 | 42.74 | 27.06
Radiographic imaging QA | The image quality degraded, resulting in lower matching accuracy. | 4.67 | 2.34 | 6.33 | 1.86 | 5.2 | 0.54772 | 148.42 | 83.89 | 2.29 | 0.95 | 3.86 | 1.21 | 5.27 | 0.53452 | 49.99 | 31.39
Radiographic imaging QA | The QC phantom measured with incorrect energy value. | 3.33 | 0.52 | 4.67 | 1.37 | 5.85 | 0.16432 | 88.9 | 20.76 | 2.86 | 0.69 | 3.71 | 0.49 | 5.87 | 0.16036 | 62.91 | 19.75
Radiographic imaging QA | The QC phantom was measured at the wrong angle. | 3 | 0.63 | 4.67 | 1.75 | 5.75 | 0.49295 | 81.98 | 38.6 | 2.57 | 0.79 | 4 | 0.82 | 5.81 | 0.48107 | 60.43 | 25.27
Radiographic imaging QA | Measured at the wrong distance. | 3.17 | 1.17 | 4.17 | 0.75 | 5.57 | 0.46762 | 72.85 | 31.35 | 2.57 | 0.98 | 4 | 0.82 | 5.41 | 0.1069 | 57.4 | 29.9
QA Data analyze | Out-of-tolerance measurements were overlooked. | 4 | 1.55 | 4.83 | 1.6 | 7.47 | 0.8165 | 144.17 | 72.91 | 2.71 | 0.95 | 3.86 | 0.38 | 7.8 | 9.5934e-16 | 83.57 | 32.07
QA Data analyze | Incorrect setting of analysis software. | 3.5 | 0.84 | 4 | 0 | 6.17 | 0.62823 | 86.2 | 21.71 | 2.86 | 0.9 | 3.86 | 0.69 | 6.04 | 0.32071 | 64.97 | 17.22
QA Data analyze | The images to be analyzed were captured at angles other than the original gantry angle. | 3.33 | 1.03 | 4.67 | 1.63 | 5.25 | 0.49295 | 86.4 | 53.41 | 2.14 | 1.07 | 3.57 | 0.53 | 5.31 | 0.48107 | 40.37 | 19.15
QA Data analyze | Different data were compared. | 3.33 | 0.52 | 4 | 0.63 | 4.95 | 0.27386 | 66.25 | 15.43 | 2.71 | 0.95 | 3.86 | 1.07 | 4.99 | 0.26726 | 51.34 | 18.56
QA Data analyze | Image resolution was beyond the specifications of the analyzer. | 3.17 | 0.75 | 5 | 2.45 | 5 | 0 | 73.33 | 16.33 | 2.29 | 0.95 | 3.71 | 0.95 | 5 | 0 | 45.71 | 28.2
QA Data analyze | Manual was not understood, and QC tools were prepared and analyzed incorrectly. | 4.83 | 1.47 | 5 | 1.55 | 5.35 | 0.38341 | 133.72 | 74.6 | 4.86 | 1.77 | 3.86 | 0.69 | 5.4 | 0.37417 | 99.86 | 34.95